## Tests (CARB Specific)
Standard Test Cycles (US) | CARB Naming
EPA 75, 3-Bag | EC1
EPA Highway Fuel Economy (HWFET) | HWY15
EPA US06 | US1B
EPA SC03 | SC03
Traces | CARB Naming
EPA75 P1 | Horiba defaults can be used
EPA75 P2 | Horiba defaults can be used
HWFET | Horiba defaults can be used
US06 | Horiba defaults can be used
SC03 | Horiba defaults can be used
Shift Lists | CARB Naming
EPA 75 3 Bag AT | Horiba defaults can be used
HWFET AT | Horiba defaults can be used
US06 AT | Horiba defaults can be used
SC03 AT | Horiba defaults can be used
Prep Cycles | CARB Naming
EPA75 P1
EPA75 P2
HWFET
US06
SC03
Naming Options
Test procedure has the same name as in the VTS system
Test procedure in STARS VETS has standard name and will be mapped to VTS name. VTS name will be copied into a customized channel that can be added to the report
| Regulation |  |  |  | CFR86_FTP
| Speed Units |  |  |  | X
RCB correction criterion (Fraction)
Downscaling method
EPA 75 3 Bag
Drivers Aid tolerance bands visible
Drivers Aid target speed visible
Fan control mode
Drivers Aid minimum time - relative (s)
Drivers Aid maximum time - relative (s)
| Blocks | Name
Speed tolerance (mph;km/h)
Time tolerance (s)
Speed out-of-tolerance time (s)
|  | Trace start on | Pendant Start
Pendant Run
|  | Block end on | Trace Finish
Pendant Stop
Drivers Aid minimum speed
Drivers Aid maximum speed
Traces
Events
Name
Minimum duration (s)
Maximum duration (s)
Name
Speed tolerance (mph;km/h)
Time tolerance (s)
Speed out-of-tolerance time (s)
|  | Trace start on | Pendant Start
Pendant Run
|  | Block end on | Trace Finish
Pendant Stop
Drivers Aid minimum speed
Drivers Aid maximum speed
Traces
Events
Name
Minimum number of coastdowns
Maximum number of coastdowns
|  | Measured data process method | No smoothing
Smoothing method 1
|  | Judge criteria | Force
Fraction (Force)
Fraction (Time)
|  | Speed Range | Row(s) | Speed upper
Speed lower
Error limit force (N)
Error limit fraction (%)
Name
Measurement time (s)
Minimum number of idle checks
Maximum number of idle checks
Calibrate analyzers
Purge analyzers
|  | IDLE check type | Low
Low and High
IDLE high minimum value (rpm)
Note on A39: Siebler, Jorg: Select one of the options listed below

## Tests
EPA 75 3 Bag | Regulation |  |  |  |  | CFR1066
| Speed Units | Miles per hour |  |  |  | X
|  | Kilometers per hour |  |  |  | -
| RCB correction criterion (Fraction) |  |  |  |  | -
| Downscaling method | WLTP class 1 |  |  |  | -
|  | WLTP class 2 |  |  |  | -
|  | WLTP class 3 |  |  |  | -
| Drivers Aid tolerance bands visible |  |  |  |  | X
| Drivers Aid target speed visible |  |  |  |  | X
| Fan control mode | Proportional |  |  |  | -
|  | Constant Speed |  |  |  | X
| Drivers Aid minimum time - relative (s) |  |  |  |  | -10
| Drivers Aid maximum time - relative (s) |  |  |  |  | 20
| Blocks | Drive Unit | Name |  |  | DriveUnit
|  |  | Speed tolerance (mph;km/h) |  |  | 2
|  |  | Time tolerance (s) |  |  | 1
|  |  | Speed out-of-tolerance time (s) |  |  | 2
|  |  | Trace start on | Pendant Start |  | -
|  |  |  | Pendant Run |  | X
|  |  | Block end on | Trace Finish |  | X
|  |  |  | Pendant Stop |  | -
|  |  | Drivers Aid minimum speed |  |  | 0
|  |  | Drivers Aid maximum speed |  |  | 60
|  |  | Traces |  |  | EPA75 P1 EPA75 P2
Events
|  | Soak | Name |  |  | Soak
|  |  | Minimum duration (s) |  |  | 545
|  |  | Maximum duration (s) |  |  | 605
|  | Drive Unit | Name |  |  | DriveUnit1
|  |  | Speed tolerance (mph;km/h) |  |  | 2
|  |  | Time tolerance (s) |  |  | 1
|  |  | Speed out-of-tolerance time (s) |  |  | 2
|  |  | Trace start on | Pendant Start |  | -
|  |  |  | Pendant Run |  | X
|  |  | Block end on | Trace Finish |  | X
|  |  |  | Pendant Stop |  | -
|  |  | Drivers Aid minimum speed |  |  | 0
|  |  | Drivers Aid maximum speed |  |  | 60
|  |  | Traces |  |  | EPA75 P1
|  | Coast Down | Name |  |  | CoastDown
|  |  | Minimum number of coastdowns |  |  | 1
|  |  | Maximum number of coastdowns |  |  | 3
|  |  | Measured data process method | No smoothing |  | X
|  |  |  | Smoothing method 1 |  | -
|  |  | Judge criteria | Force |  | X
|  |  |  | Fraction (Force) |  | -
|  |  |  | Fraction (Time) |  | -
|  |  | Speed Range | Row(s) | Speed upper | 65
[5 empty] | Speed lower | 55
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 55
[5 empty] | Speed lower | 45
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 45
[5 empty] | Speed lower | 35
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 35
[5 empty] | Speed lower | 25
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 25
[5 empty] | Speed lower | 15
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 15
[5 empty] | Speed lower | 5
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  | IDLE Check | Name
Measurement time (s)
Minimum number of idle checks
Maximum number of idle checks
Calibrate analyzers
Purge analyzers
|  |  | IDLE check type | Low
Low and High
IDLE high minimum value (rpm)
HWFET | Regulation |  |  |  |  | CFR1066
| Speed Units | Miles per hour |  |  |  | X
|  | Kilometers per hour |  |  |  | -
| RCB correction criterion (Fraction) |  |  |  |  | 0
| Downscaling method | WLTP class 1 |  |  |  | -
|  | WLTP class 2 |  |  |  | -
|  | WLTP class 3 |  |  |  | -
| Drivers Aid tolerance bands visible |  |  |  |  | X
| Drivers Aid target speed visible |  |  |  |  | X
| Fan control mode | Proportional |  |  |  | -
|  | Constant Speed |  |  |  | X
| Drivers Aid minimum time - relative (s) |  |  |  |  | -10
| Drivers Aid maximum time - relative (s) |  |  |  |  | 20
| Blocks | Drive Unit | Name |  |  | DriveUnit
|  |  | Speed tolerance (mph;km/h) |  |  | 2
|  |  | Time tolerance (s) |  |  | 1
|  |  | Speed out-of-tolerance time (s) |  |  | 2
|  |  | Trace start on | Pendant Start |  | -
|  |  |  | Pendant Run |  | X
|  |  | Block end on | Trace Finish |  | X
|  |  |  | Pendant Stop |  | -
|  |  | Drivers Aid minimum speed |  |  | 0
|  |  | Drivers Aid maximum speed |  |  | 80
|  |  | Traces |  |  | HWFET
|  | Coast Down | Name |  |  | CoastDown
|  |  | Minimum number of coastdowns |  |  | 1
|  |  | Maximum number of coastdowns |  |  | 3
|  |  | Measured data process method | No smoothing |  | X
|  |  |  | Smoothing method 1 |  | -
|  |  | Judge criteria | Force |  | X
|  |  |  | Fraction (Force) |  | -
|  |  |  | Fraction (Time) |  | -
|  |  | Speed Range | Row(s) | Speed upper | 65
[5 empty] | Speed lower | 55
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 55
[5 empty] | Speed lower | 45
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 45
[5 empty] | Speed lower | 35
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 35
[5 empty] | Speed lower | 25
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 25
[5 empty] | Speed lower | 15
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 15
[5 empty] | Speed lower | 5
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
US06 | Regulation |  |  |  |  | CFR1066
| Speed Units | Miles per hour |  |  |  | X
|  | Kilometers per hour |  |  |  | -
| RCB correction criterion (Fraction) |  |  |  |  | 0
| Downscaling method | WLTP class 1 |  |  |  | -
|  | WLTP class 2 |  |  |  | -
|  | WLTP class 3 |  |  |  | -
| Drivers Aid tolerance bands visible |  |  |  |  | X
| Drivers Aid target speed visible |  |  |  |  | X
| Fan control mode | Proportional |  |  |  | -
|  | Constant Speed |  |  |  | X
| Drivers Aid minimum time - relative (s) |  |  |  |  | -10
| Drivers Aid maximum time - relative (s) |  |  |  |  | 20
| Blocks | Drive Unit | Name |  |  | DriveUnit
|  |  | Speed tolerance (mph;km/h) |  |  | 2
|  |  | Time tolerance (s) |  |  | 1
|  |  | Speed out-of-tolerance time (s) |  |  | 2
|  |  | Trace start on | Pendant Start |  | -
|  |  |  | Pendant Run |  | X
|  |  | Block end on | Trace Finish |  | X
|  |  |  | Pendant Stop |  | -
|  |  | Drivers Aid minimum speed |  |  | 0
|  |  | Drivers Aid maximum speed |  |  | 100
|  |  | Traces |  |  | US06
|  | Coast Down | Name |  |  | CoastDown
|  |  | Minimum number of coastdowns |  |  | 1
|  |  | Maximum number of coastdowns |  |  | 3
|  |  | Measured data process method | No smoothing |  | X
|  |  |  | Smoothing method 1 |  | -
|  |  | Judge criteria | Force |  | X
|  |  |  | Fraction (Force) |  | -
|  |  |  | Fraction (Time) |  | -
|  |  | Speed Range | Row(s) | Speed upper | 65
[5 empty] | Speed lower | 55
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 55
[5 empty] | Speed lower | 45
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 45
[5 empty] | Speed lower | 35
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 35
[5 empty] | Speed lower | 25
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 25
[5 empty] | Speed lower | 15
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 15
[5 empty] | Speed lower | 5
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
SC03 | Regulation |  |  |  |  | CFR1066
| Speed Units | Miles per hour |  |  |  | X
|  | Kilometers per hour |  |  |  | -
| RCB correction criterion (Fraction) |  |  |  |  | 0
| Downscaling method | WLTP class 1 |  |  |  | -
|  | WLTP class 2 |  |  |  | -
|  | WLTP class 3 |  |  |  | -
| Drivers Aid tolerance bands visible |  |  |  |  | X
| Drivers Aid target speed visible |  |  |  |  | X
| Fan control mode | Proportional |  |  |  | -
|  | Constant Speed |  |  |  | X
| Drivers Aid minimum time - relative (s) |  |  |  |  | -10
| Drivers Aid maximum time - relative (s) |  |  |  |  | 20
| Blocks | Drive Unit | Name |  |  | DriveUnit
|  |  | Speed tolerance (mph;km/h) |  |  | 2
|  |  | Time tolerance (s) |  |  | 1
|  |  | Speed out-of-tolerance time (s) |  |  | 2
|  |  | Trace start on | Pendant Start |  | -
|  |  |  | Pendant Run |  | X
|  |  | Block end on | Trace Finish |  | X
|  |  |  | Pendant Stop |  | -
|  |  | Drivers Aid minimum speed |  |  | 0
|  |  | Drivers Aid maximum speed |  |  | 80
|  |  | Traces |  |  | SC03
|  | Coast Down | Name |  |  | CoastDown
|  |  | Minimum number of coastdowns |  |  | 1
|  |  | Maximum number of coastdowns |  |  | 3
|  |  | Measured data process method | No smoothing |  | X
|  |  |  | Smoothing method 1 |  | -
|  |  | Judge criteria | Force |  | X
|  |  |  | Fraction (Force) |  | -
|  |  |  | Fraction (Time) |  | -
|  |  | Speed Range | Row(s) | Speed upper | 65
[5 empty] | Speed lower | 55
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 55
[5 empty] | Speed lower | 45
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 45
[5 empty] | Speed lower | 35
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 35
[5 empty] | Speed lower | 25
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 25
[5 empty] | Speed lower | 15
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
|  |  |  | Row(s) | Speed upper | 15
[5 empty] | Speed lower | 5
[5 empty] | Error limit force (N) | 10
[5 empty] | Error limit fraction (%) | 0
Note on B3: Siebler, Jorg: Available for demo: CFR1066 CFR86 Additional in Standard: Euro3 Euro4 Euro5 Euro6 (b/c) Japan 1015 Japan JC08
Note on D16: Siebler, Jorg: a user-defined name for the Drive Unit Block (this is the name that is displayed in the Drive Unit Block header)
Note on D17: Siebler, Jorg: the maximum value by which the driver is allowed to deviate from any speed setpoint
Note on D18: Siebler, Jorg: the maximum time the driver is allowed to deviate from any time setpoint
Note on B85: Siebler, Jorg: Available for demo: CFR1066 CFR86 Additional in Standard: Euro3 Euro4 Euro5 Euro6 (b/c) Japan 1015 Japan JC08
Note on B143: Siebler, Jorg: Available for demo: CFR1066 CFR86 Additional in Standard: Euro3 Euro4 Euro5 Euro6 (b/c) Japan 1015 Japan JC08
Note on B201: Siebler, Jorg: Available for demo: CFR1066 CFR86 Additional in Standard: Euro3 Euro4 Euro5 Euro6 (b/c) Japan 1015 Japan JC08

## Resource Mapping Der (CARB XML)
|  | Label | Channel | Add to report | Report Section
Vehicles | Name |  | "VTS " + TEST_ID_NUM | X | VehicleDetails
| Road load model identifier |  | N/A
| Units for vehicle model | SI | TARGET_COEFF_UNIT
N-(km/h)
US
| Highway | Inertia | HIGHWAY_INERTIA_WEIGHT | X | VehicleDetails
|  | Road load coefficient A | TARGET_COEFF_A | X | VehicleDetails
|  | Road load coefficient B | TARGET_COEFF_B | X | VehicleDetails
|  | Road load coefficient C | TARGET_COEFF_C | X | VehicleDetails
| Dyno | Inertia | ACTUAL_INERTIA_WEIGHT | X | VehicleDetails
|  | Road load coefficient A | DYNO_COEFF_A | X | VehicleDetails
|  | Road load coefficient B | DYNO_COEFF_B | X | VehicleDetails
|  | Road load coefficient C | DYNO_COEFF_C | X | VehicleDetails
| Wheelbase |  | WHEEL_BASE | X | VehicleDetails
| Roll configuration | Front wheel drive | DRIVE_MODE | X | VehicleDetails
|  | Rear wheel drive |  |  | VehicleDetails
|  | All wheel drive |  |  | VehicleDetails
| WLTP vehicle class | Class 1 | N/A
|  | Class 2 | N/A
|  | Class 3a | N/A
|  | Class 3b | N/A
| Pollutant limits | [List of configured Pollutant Limits] | Default tbd
| Powertrain | [List of configured Powertrains] | N/A
| CANdb configuration | [List of configured CAN Configurations] | N/A
| Custom Fields | Vehicle ID | VEHICLE_ID | X | VehicleDetails
|  | VIN | VIN | X | VehicleDetails
|  | Odometer | ODOMETER | X | VehicleDetails
|  | Model Year | MODEL_YEAR | X | VehicleDetails
|  | Modal Type | MODEL_TYPE | X | VehicleDetails
|  | Manufacturer Name | MFR_NAME | X | VehicleDetails
|  | Manufacturer Division Name | DIVISION_NAME | X | VehicleDetails
|  | Engine Family Number | ENGINE_FAMILY | X | VehicleDetails
|  | Engine Displacement (ltr) | ENGINE_DISP_VOL_LITRE | X | VehicleDetails
|  | Transmission Type | TRANS_TYPE_CODE | X | VehicleDetails
|  | Number of Cylinders | CYLINDER_NUM | X | VehicleDetails
|  | License Number | LICENSE_NUM | X | VehicleDetails
|  | State Code | STATE | X | VehicleDetails
|  | Vehicle Class Code | VEH_CLASS_CODE | X | VehicleDetails
|  | EO Inertia | EO_INERTIA_WEIGHT | X | VehicleDetails
|  | EO Horsepower | EO_HORSEPOWER | X | VehicleDetails
|  | EO Horsepower Flag | EO_HORSEPOWER_FLAG | X | VehicleDetails
|  | Horsepower Modifier | ADD_10_PERCENT | X | VehicleDetails
|  | Road Load Coefficient D (rad) | TARGET_COEFF_D | X | VehicleDetails
|  | Target Level Unit | TARGET_LEVEL_UNIT | X | VehicleDetails
|  | Road Load Coefficient N (lbs) | TARGET_COEFF_N | X | VehicleDetails
|  | Target Level Type | TARGET_LEVEL_TYPE | X | VehicleDetails
|  | Horsepower Actual Value | HP_ACTUAL_VALUE | X | VehicleDetails
|  | Mileage Accumulation | MILEAGE_ACCUMULATION | X | VehicleDetails
|  | Dyno Coefficient Set Number | DYNO_COEFF_SET_NUMBER | X | VehicleDetails
|  | Dyno Coefficient D | DYNO_COEFF_D | X | VehicleDetails
|  | Dyno Coefficient N | DYNO_COEFF_N | X | VehicleDetails
|  | Dyno Coefficient Unit | DYNO_COEFF_UNIT | X | VehicleDetails
|  | Tire Type | TIRE_TYPE | X | VehicleDetails
|  | Tire Size | TIRE_SIZE | X | VehicleDetails
|  | Actual Tire Presure | ACTUAL_TIRE_PRESSURE | X | VehicleDetails
|  | Equivalent Inertia Mass (kg) | EQUIV_INERTIA_MASS | X | VehicleDetails
|  | Road Load Force (N) | ROAD_LOAD_FORCE | X | VehicleDetails
|  | Odometer Units | ODOMETER_UNIT_CODE | X | VehicleDetails
|  | Dyno Direction | DYNO_DIRECTION | X | VehicleDetails
|  | Dyno Synch | DYNO_SYNCH | X | VehicleDetails
|  | Drive Wheels Type Code | DRIVE_WHEELS_TYPE_CODE | X | VehicleDetails
Fuels | Name |  | "VTS " + TEST_ID_NUM | X | FuelInfo
| Fuel Type | [List of configured Fuel Types in Custom Enumerations] | FUEL_TYPE | X | FuelInfo
| H to C ratio |  | H_C_RATIO | X | FuelInfo
| O to C ratio |  | N/A
| S to C ratio |  | N/A
| N to C ratio |  | N/A
| Non methane HC H to C ratio |  | NMHC_CWF | X | FuelInfo
| Fuel specific gravity |  | SPECIFIC_GRAVITY_60F | X | FuelInfo
| Fuel lower heating value |  | NHV | X | FuelInfo
| R factor |  | N/A
| H mass fraction |  | N/A
| C mass fraction |  | CWF | X | FuelInfo
| Custom Fields | Natural Gas CO2 Mass Fraction (%Mass) | CO2_NG | X | FuelInfo
|  | Fuel HC Carbon Mass Fraction (%Mass) | HC_CWF | X | FuelInfo
|  | Natural Gas Density | DENSITY_NG | X | FuelInfo
|  | Gasoline Equivalent Conversion Factor | GECF | X | FuelInfo
|  | Dilution Factor Code | DILUTION_FACTOR_CODE | X | FuelInfo
|  | F2 | F2 | X | FuelInfo
|  | F3 | F3 | X | FuelInfo
|  | F4 | F4 | X | FuelInfo
|  | F4A | F4A | X | FuelInfo
|  | Sample ID | SAMPLE_ID | X | FuelInfo
|  | Batch ID | BATCH_ID | X | FuelInfo
Tests | Name |  | "VTS " + TEST_ID_NUM | X | TestReference
| Settings | Test Procedure | TEST_NAME (needs to be alligned) | X | TestInformation
|  | Shift List | TEST_NAME (needs to be alligned) | X | TestInformation
|  | Configuration | TEST_NAME (needs to be alligned) | X | TestInformation
|  | Sample Line Configuration | TEST_NAME (needs to be alligned) | X | TestInformation
|  | Regulation | Default tbd | X | TestInformation
|  | Fuel | "VTS " + TEST_ID_NUM | X | TestInformation
|  | Vehicle | "VTS " + TEST_ID_NUM | X | TestInformation
|  | Channel Set | TEST_NAME (needs to be alligned)
|  | SAO Offset | Standard
?
|  | Run export after test | Default tbd
|  | Run report after test | Default tbd
|  | Run idle check | Default tbd
|  | Idle check type | Default
?
|  | Idle high minimum value (rpm) | Default tbd
|  | Prompt to repeat test procedure | Default tbd
|  | Wait for bag reads | Default tbd
|  | Bag pair dump duration (s) | Default tbd
|  | Run vacuum side leak check | Default tbd
|  | Run CVS blower | Default tbd
|  | Control chassis dynamometer | Default tbd
|  | Run coastdown | Default tbd
|  | Control gradient demand | Default tbd
| Custom Fields | VTS Status | TEST_STATUS
|  | VTS Test Name | TEST_NAME | X | TestInformation
|  | Test ID Number | TEST_ID_NUM | X | TestInformation
|  | Project ID | PROJECT_ID | X | TestInformation
|  | Vehicle Project Sequence Number | VEH_PROJ_SEQ_NUM | X | TestInformation
|  | Test Type Code | TEST_TYPE_CODE | X | TestInformation
|  | Driver ID | DRIVER_ID | X | TestInformation
|  | Operator ID | OPERATOR_ID | X | TestInformation
|  | Supervisor ID | SUPERVISOR_ID | X | TestInformation
|  | Violation Schedule File | VIOLATION_SCHEDULE_FILE
|  | Augmented Braking Schedule File | AUG_BRAKING_SCHEDULE_FILE
|  | Aldehyde Sampling Schedule File | ALDEHYDE_SCHEDULE_FILE
|  | Auxiliary Analog Output Schedule File | AUX_ANALOG_SCHEDULE_FILE
|  | CVS VVS Schedule File | CVS_VVS_SCHEDULE_FILE
|  | Test Cell | TEST_CELL | X | TestInformation
|  | Suggested Test Cell | SUGGESTED_TEST_CELL
|  | Priority | PRIORITY
|  | Procedure Code | PROCEDURE_CODE
|  | Test Stand Code | TEST_STAND_CODE | X | TestInformation
|  | Comments | COMMENTS | X | TestInformation
|  | Modification Date | MODIFICATION_DATE
|  | Modified By | MODIFIED_BY
|  | Test Date | TEST_DATE
|  | Used Robot Driver | USED_ROBOT_DRIVER
|  | Grade Schedule File | GRADE_SCHEDULE_FILE
|  | ProjectVehicleTestTypeSequenceNumber | PROJ_VEH_TEST_TYPE_SEQ_NUM
VTS Fields
Unclear | PROJ_DIAG_REP_SEQ_NUM | Diagnostic & Repair sequence number within a project. This is a sequential number used to keep track of repair sequences for the vehicle within a project
| DIAG_REPAIR_REASON_CODE | Diagnostic & Repair reason code. The code represents the reason for the diagnostic & repair sequence number.
CANISTER_CONDITIONING
| BED_VOLUME | Not used
| WORKING_CAPACITY | Not used
| LOAD_FLOW_RATE | Not used
| PREP | Not used
| LOW_TEMP | Not used
| HIGH_TEMP | Not used
| SOAK_TIME_MIN | Not used
| SOAK_TIME_MAX | Not used
| BACKPRESSURE | Not used
| AUX_DIGITAL_SCHEDULE_FILE | Not used
| PRECON_TYPE_CODE | The precondition code that represents all preconditioning task.
| CVS_FLOW_RATE | Not used (this is generally set at the test cell by the operator
| DUAL_BENCH_CLEAN | Flag to designate if the test will be using dual bench configuration with the clean bench
| RL_TEMP | Not used (pertains to Running Loss)
| SHIFT_SCHEDULE_FILE | Not used
| RL_TEST_TYPE | Not used (pertains to Running Loss)
| WHEEL_BASE -> What are the units? Inches? Feet? | inches
| TARGET_COEFF_UNIT & DYNO_COEFF_UNIT -> What are the possible values aside from 'ENGLISH'. If there are other options, what do they mean (i.e. ENGLISH -> US) | SI is the other option (metric)
| DYNO_COEFF_UNIT -> If the entry has no value what are theunits of DYNO_COEFF_A, DYNO_COEFF_B, DYNO_COEFF_C | Default should use ENGLISH units
| TEST_TYPE_CODE -> What are possible codes and corresponding TEST_NAME values | See the entries on the "Tests (CARB Specific) tab of this spreadsheet
| FUEL_TYPE -> What are all possible codes and meanings of each code name | See added tab "ARB FuelDesciption"
| VEH_CLASS_CODE -> What are all possible codes and meanings of each code name | See added tab "ARB Veh_Class_Code"
Flags | DOUBLE_SAMPLE_FLAG
SPECIATION_FLAG
| ALDEHYDE_FLAG | Impinger
| ALCOHOL_FLAG | Impinger
SET_BACKPRESSURE_FLAG
| N2O_SAMPLE_FLAG | MEXA-1100QL NO2
| PARTICULATE_SIZER_FLAG | SPCS2000
| PARTICULATE_SPECIATION_FLAG | TSI
| SOOT_FLAG | AVL483
| PARTICULATE_FLAG | Quad PM
HFID_FLAG
POWER_METER_FLAG
Note on D130: Siebler, Jorg: Always Cell #1
Note on D131: Siebler, Jorg: Can we list in STARS VETS based on the priority?

## ARB FuelDescription
FUEL_TYPE | DILUTION_FACTOR_CODE (G=Gasoline, L=LPG, A=Alcohol, N = CNG, P=Phase 2 Gasoline, D=Diesel) | FUEL_NAME | FUEL_COMMENT
BN01 | A | BUTANOL 100%
BN02 | A | BUTANOL 85% + 15% INDOLENE CLEAR
DC01 | G | DIESEL 25% SULFUR [CERT FUEL PER MY94]
DC02 | G | DIESEL 10% SULFUR [OLD EPA, MY90-94]
DC03 | G | DIESEL 05% SULFUR [40 CFR 86-1313-94 (b)]
DC04 | G | CA CLEAN DIESEL FUEL AKA DIESEL 13 CERT FUEL STARTING IN MY95
DC05 | D | California Cert Diesel
DC06 | D | Cert Diesel analysis data from vendor
DC07 | D | Fuel for AIGER crosscheck
DC08 | D | Cert Fuel Diesel analysis from Paragon Laboratories
DN01 | G | DIESEL COMMERCIAL #1
DN02 | D | DIESEL COMMERCIAL #2
DN03 | G | RENTECH DIESEL
DN04 | G | DIESEL PRE 1985
DN05 | G | CA Diesel - MSOD 4R08C01
DN06 | G | B100 Animal Diesel | 2/26/2009 9:5 2R08C01 3/3/2009 8:18 Revised project # 4R08C01
DN07 | G | B100 Soy Diesel | 2/26/2009 9:7 2R08C01 3/3/2009 8:37 Revised project # 4R08C01
DN08 | G | B50 Soy Diesel | 2/26/2009 9:9 2R08C01 3/3/2009 8:37 Revised project # 4R08C01
DN09 | G | B20 Soy Diesel | 2/26/2009 9:10 2R08C01 3/3/2009 8:38 Revised project # 4R08C01
DN10 | G | B5 Soy Diesel | 2/26/2009 9:11 2R08C01 3/3/2009 8:38 Revised project # 4R08C01
DN11 | G | B-99 S+
DN12 | D | Soy diesel | 07/19/2012 11:14 New Soy (20%) diesel 07/24/2012 09:26 Update - percent bio is being researched.
E10 | G | CARB LEE III E10 Certification Fuel, Halterman Product Cod: HF0862 Batch DJ0921L
EC01 | A | ETHANOL 98% + 2% INDOLENE CLEAR
EC02 | A | ETHANOL 85% + 15% CERTIFICATION PHASE 2 GASOLINE | 4/12/2002 16:2 Fuel supplied by US Liquides at Ed Sun's request
EC03 | A | 85% Ethanol + 15% gasoline | 6/5/2002 11:13 Future Truck testing fuel
EC04 | G | E10 Specialty Fuel CE10 For Marine Projects | 10/2/2007 10:16 45% Toluene, 45% Isooctane and 10% Ethanol
EC05 | A | E-85 test fuel | Per Meng: Please change MPG Calculation from G to A (For Ethanol fuel)
EC06 | G | E-10 Cert fuel Phase III
EC07 | G | E-85 (EPA Tier II EEE Cert. fuel)
EC08 | G | E-10 Premium Cert.
EC09 | G | Tank 1 E-10 Cert
EC10 | G | CARB LEV III E10 Certification Fuel Premium Octane, Halterman Product Code: HF08
EN01 | A | ETHANOL 100%
EN02 | A | ETHANOL 95% + 5% INDOLENE CLEAR
EN03 | A | ETHANOL 85% + 15% RE-FORMULATED RF-A GASOLINE
EN04 | A | ETHANOL 85% + 15% INDOLENE CLEAR
EN05 | A | E85-Possibel Cert. E85 Fuel | 10/25/2006 10:28 E85-Possible Cert E85 speciay made by Chevron for E85 demo project amd used for porject 2R0601 (SFTP)
EN06 | A | Chevron E-85 fuel for project 2R0601
EN07 | G | Chevron E-10 fuel
EN08 | G | E-10 fuel used in 2R0705
EN09 | G | E-10 for project 2R0911
EN10 | G | E-85 from test vehicle 107
EN11 | G | Commercial E-85 fuel
EV00 | G | Fuel for EV tests
GC01 | P | CERTIFICATION PHASE 2 GASOLINE
GC02 | P | CERTIFICATION PHASE 2 GASOLINE [MY92-94]
GC03 | G | Cert fuel project# 2R1108
GC04 | P | Reformulated Gasoline | 6/5/2002 11:12 Fuel for Future Truck Testing
GC05 | G | Federal Tier 2 Cert
GC06 | G | EPA Tier 3 EEE Certification
GK01 | P | phase 2 certification fuel
GN01 | P | COMMERCIAL PHASE 2 GASOLINE
GN02 | G | FEDERAL PHASE 1 GASOLINE
GN03 | P | GASOLINE PHASE 2 90% HC + 10% ETHANOL
GN04 | P | GASOLINE PHASE 2 89% HC + 11% MTBE
GN05 | G | REFORMULATED RF-A GASOLINE [PHILLIPS 66]
GN06 | G | METHYL TERTIARY BUTYL ETHER [MTBE]
GN07 | G | GASOHOL [COMMERCIAL GRADE]
GN08 | G | LEADED COMMERCIAL TANK FUEL
GN09 | G | FED SUMMER RFG [FES0001]
GN10 | G | CA PHASE 1 [FES0003]
GN11 | P | COMMERCIAL PHASE 2 GASOLINE USED IN PRE-PHASE TESTING
GN12 | G | GM 92 PRE-PHASE 1 SUMMER FUEL (DRUM)
GN13 | G | GM 92 PRE-PHASE 1 WINTER FUEL (DRUM)
GN14 | G | CHEVRON NON-OXYGENATED PHASE 2 FUEL (DURM)
GN15 | G | TOSCO 2% OXYGENATED ETHANOL FUEL (DRUM)
GN16 | P | COMMERCIAL PHASE 2 WINTER GRADE, tank 2
GN17 | G | Commercial Gasoline | 11/6/2002 13:20 Kleeman PM project
GN18 | P | Commercial gasoline | 6/17/2003 13:5 CE-CERT ARB NMOG Correlation study
GN19 | G | Commercial Phase 3 gasoline
GN20 | G | winter commercial phase 3 reformulated gasoline
GN21 | G | Commercial 91 octane
GN22 | G | Commercial Phase 3 Gasoline - Regular grade | Purchased for local station for project 2R0605
GN23 | G | Commercial Phase 3 Gasoline - Premium grade | Purchased for local station for project 2R0605
GN24 | G | Commercial 91 2R0406 | 10/25/2006 10:41 OBD Project 2R0406
GN25 | G | Commercial Regular (Jetta GHG Project)
GN26 | G | 10% ethanol fuel
GN27 | G | Commercial fuel (E6) from 76 gas station for project 2R0905
GN28 | G | Multiple test fuels for comparison testing from project 2R0905 | 6/2/2009 8:24 This fuel code should not be used for any other testing. 6/2/2009 8:24 This fuel code should not be used for any other testing. 6/3/2009 11:49 PROJECT 2R0905 ONLY
HN01 | G | HYDROGEN
IC01 | G | Indolene Certification Gasoline
IC02 | G | CAA baseline indolene clear certification fuel
IC21 | G | Indolene Clear Certification Fuel Tier 2 (unleaded)
IC22 | G | EPA unleaded Tier 2 | 9/9/2005 13:11 EPA fuel type 61
IC23 | G | Tier 1 Indolene
IN01 | G | INDOLENE 30
IN02 | G | INDOLENE 15
IN03 | G | INDOLENE CLEAR 90% + 10% ETHANOL
IN04 | G | INDOLENE CLEAR 75% + 25% METHANOL
IN05 | G | INDOLENE CLEAR 65% + 35% METHANOL
IN06 | G | INDOLENE 53% + 47% WATER
LC01 | L | LIQUIFIED PETROLEUM GAS -- CERTIFICATION (6/98, BATCH 002)
LC02 | L | LPG SPECIAL TEST: CERT. FUEL, 94% PROPANE, 4% PROPYLENE, 2% BUTANE
LN01 | L | LIQUIFIED PETROLEUM GAS -- COMMERCIAL
LN02 | L | LIQUIFIED PETROLEUM GAS -- UNKNOWN
LN03 | L | LIQUIFIED PETROLEUM GAS -- TEST FUEL (6/98, BATCH 003)
LN04 | L | LIQUIFIED PETROLEUM GAS -- TEST FUEL (6/98, BATCH 004)
LN05 | L | LIQUIFIED PETROLEUM GAS -- TEST FUEL (6/98, BATCH 005)
LN06 | L | LIQUIFIED PETROLEUM GAS -- TEST FUEL (6/98, BATCH 006)
LN07 | L | LIQUIFIED PETROLEUM GAS -- TEST FUEL (8/98, BATCH 007)
LN08 | L | LIQUIFIED PETROLEUM GAS -- TEST FUEL (8/98, BATCH 008)
LN09 | L | LIQUIFIED PETROLEUM GAS from vehicle
MC01 | A | METHANOL 100% CERTIFICATION FUEL
MC02 | A | METHANOL 85% + 15% APPLICABLE CERT GASOLINE
MC03 | A | METHANOL 85% + 15% PHASE 2 CERTIFICATION GASOLINE
MN01 | A | METHANOL 95% + 5% INDOLENE CLEAR
MN02 | A | METHANOL 90% + 10% COMMERCIAL UNLEADED GASOLINE
MN03 | A | METHANOL 88% + 12% COMMERCIAL UNLEADED GASOLINE
MN04 | A | METHANOL 85% + 15% UNKNOWN GASOLINE
MN05 | A | METHANOL 85% + 15% REFORMULATED RF-A GASOLINE
MN06 | A | METHANOL 50% + 50% INDOLENE CLEAR
NC01 | N | COMPRESSED NATURAL GAS -- CERTIFICATION FUEL
NN01 | N | COMPRESSED NATURAL GAS -- COMMERCIAL FUEL
NN02 | N | Commercial Compressed Natural Gas
NN03 | N | Commercial CNG Fuel | 2/7/2007 7:54 Commercial CNG Fuel purchased from SoCalGas
NN04 | N | LIQUIFIED NATURAL GAS
NOFL | G | No Fuel for SHED Test
TT8D | N | Test Fuel for METS Testing -- Should not be used
UN02 | G | OTHERS
VN01 | G | ELECTRIC
WN01 | G | WATER 60% + 40% GASOLINE
WN02 | A | WATER 50% + 50% ETHANOL
XN01 | G | OCTAMIX 100%
XN02 | G | OCTAMIX 7%

## ARB Veh_Class_Code
Column Name | SEQ | SHORT_VALUE | LONG_VALUE
VEH_CLASS_CODE | 1 | PC | Passenger Car
VEH_CLASS_CODE | 2 | T1 | Light Duty Truck [LDT: 0 thru 3750 lb LVW]
VEH_CLASS_CODE | 3 | T2 | Light Duty Truck [LDT: 3751 thru 5750 lb LVW]
VEH_CLASS_CODE | 4 | LDT | General Light Duty Truck
VEH_CLASS_CODE | 5 | LDV | Light Duty Vehicle [no longer used]
VEH_CLASS_CODE | 6 | MDV | Medium Duty Vehicle
VEH_CLASS_CODE | 7 | M1 | MDV [0 thru 3750 TW]
VEH_CLASS_CODE | 8 | M2 | MDV [3751 thru 5750 lb TW]
VEH_CLASS_CODE | 9 | M3 | MDV [5751 thru 8500 lb TW]
VEH_CLASS_CODE | 10 | M4 | MDV [8501 thru 10000 lb ETW, model 95]
VEH_CLASS_CODE | 11 | M5 | MDV [10001 thru 14000 lb ETW, model 95]
VEH_CLASS_CODE | 12 | HD | Heavy Duty Vehicle
VEH_CLASS_CODE | 13 | HEV | Hybrid Electric Vehicle
VEH_CLASS_CODE | 15 | ORV | Off Road Vehicle
VEH_CLASS_CODE | 18 | OFHD | Off Road Heavy Duty Engine
VEH_CLASS_CODE | 19 | ONHD | On Road Heavy Duty Engine
VEH_CLASS_CODE | 20 | ONMD | On Road Medium Duty Engine
VEH_CLASS_CODE | 21 | ONHV | On Road Heavy Duty Vehicle
VEH_CLASS_CODE | 22 | ONMV | On Road Medium Duty Vehicle
VEH_CLASS_CODE | 23 | ULG | Utility, Lawn and Garden
VEH_CLASS_CODE | 24 | MC1A | MC1A = HMC Class I-A <50 cc
VEH_CLASS_CODE | 25 | MC1B | MC1B = HMC Class I-B 50 cc - 169 cc
VEH_CLASS_CODE | 26 | MC2 | MC2 = HMC Class II 170 cc - 279 cc
VEH_CLASS_CODE | 27 | MC3 | MC3 = HMC Class III 280 cc and over
VEH_CLASS_CODE | 28 | OFMC | OFMC = Off-Road Motorcycle
VEH_CLASS_CODE | 29 | ATV | ATV = All Terrain Vehicle
VEH_CLASS_CODE | 30 | OFUV | OFUV = Off-Road Utility Vehicle
VEH_CLASS_CODE | 31 | OFSV | OFSV = Off-Road Sport Vehicle
VEH_CLASS_CODE | 32 | SC | SC = Sand Car
VEH_CLASS_CODE | 33 | T3L2 | Light Duty Truck [LDT: 6001 thru 8500 lb LVW]

## Devices
MAHA 4WD Dyno | Type |  |  | Chassis Dyno
Driver
Version
| Communication type | TCP/IP | IP Address | 10.114.162.114
|  |  | TCP Port | 7000
|  |  | UDP Port | 7100
| Communication protocol | AK |  | X
Meta
MEXA-ONE C1 | Type |  |  | Emission Bench
Driver
Version
| Communication type | TCP/IP | IP Address
TCP Port
UDP Port
| Communication protocol | AK
|  | Meta |  | X
CVS-ONE | Type |  |  | CVS
Driver
Version
| Communication type | TCP/IP | IP Address
|  |  | TCP Port | 7711
|  |  | UDP Port | 7712
| Communication protocol | AK
|  | Meta |  | X
MEXA-ONE C1-SL | Type |  |  | Emission Bench
Driver
Version
| Communication type | TCP/IP | IP Address | 10.114.162.117
TCP Port
UDP Port
| Communication protocol | AK
|  | Meta |  | X
MEXA-7200LE | Type |  |  | Emission Bench
Driver
Version
| Communication type | TCP/IP | IP Address | 10.114.162.113
TCP Port
UDP Port
| Communication protocol | AK
|  | Meta |  | X
CVS-7200 | Type |  |  | CVS
Driver
Version
| Communication type | TCP/IP | IP Address | 10.114.162.115
TCP Port
UDP Port
| Communication protocol | AK
|  | Meta |  | X
MEXA-1100QL N2O | Type |  |  | Gas Analyzer
Driver
Version
| Communication type | TCP/IP | IP Address | 10.114.162.223
TCP Port
UDP Port
| Communication protocol | AK
Meta
MEXA-1400QL NX | Type |  |  | Gas Analyzer
Driver
Version
| Communication type | TCP/IP | IP Address | 10.114.162.222
|  |  | TCP Port | 7720
|  |  | UDP Port | 7721R
| Communication protocol | AK
Meta
GC Sample System | Type |  |  | Special
Driver
Version
| Communication type | TCP/IP | IP Address | 10.114.162.118
TCP Port
UDP Port
| Communication protocol | AK |  | X
Meta
Impinger | Type |  |  | Special
Driver
Version
| Communication type | TCP/IP | IP Address | 10.114.162.118
TCP Port
UDP Port
| Communication protocol | AK |  | X
Meta
Quad Particulate Sampler | Type |  |  | PM
Driver
Version
| Communication type | TCP/IP | IP Address | 10.114.162.71
TCP Port
UDP Port
| Communication protocol | AK |  | X
Meta
MEXA-2000 SPCS | Type |  |  | PN
Driver
Version
| Communication type | TCP/IP | IP Address
TCP Port
UDP Port
| Communication protocol | AK |  | X
Meta
TSI UCPC Interface tbd | Type |  |  | Special
Driver
Version
| Communication type | TCP/IP | IP Address
TCP Port
UDP Port
| Communication protocol | AK
Meta
TSI EEPS Interface tbd | Type |  |  | Special
Driver
Version
| Communication type | TCP/IP | IP Address
TCP Port
UDP Port
| Communication protocol | AK
Meta
AVL 483 Micro Soot | Type |  |  | Special
Driver
Version
| Communication type | TCP/IP | IP Address | -
|  |  | TCP Port | -
|  |  | UDP Port | -
| Communication type | RS232 | Serial Port | COM1
|  |  | Baud rate | 19200
|  |  | Auto-connect | X
| Communication protocol | AK |  | X
Weatherstation |  |  |  | 10.114.162.103
STARS VETS |  |  |  | 10.114.162.111

## Optional Devices
SPCS 2x00 | PND1 settings | Number of valid PND1 | 1
2
3
4
5
|  | List | Number
Dilution Factor
| PND2 settings | Number of valid PND2 | 1
2
3
4
5
|  | List | Number
Dilution Factor
PDU
Bleedoff rate (l/min)
Measure stabilize time (s)
Purge time (s)
Micro II Impinger | Interface Version | 2.7
3.0
| Number of Impinger pairs | 3
4
Reference temperature (degC)
Return sample to CVS
MEXA 1400 | Raw measurement
Bleed off (l/min)
Bleed off return

## Drivers Aid
General | Type | Wireless | X
|  | Standard | -
Pendant
Integrated Drivers Aid | Dyno diameter (mm) |  | 1219.2
Dyno pulse per revolution
| Dyno speed update interval (sec) |  | 0.05
| Dyno pulse counter variable path |  | TS:VETS_IFC1_PI01Value
Maximum pulse count
Minimum pulse count
| Dyno pulse frequency variable path |  | TS:VETS_IFC1_PI02Value
Speed acquisition method
| Speed acquisition method | Use pulse count to calculate speed | -
|  | Use frequency to calculate speed | -
|  | Use speed variable to get speed | X
| Distance acquisition method | Use pulse count to calculate distance | -
|  | Use frequency to calculate distance | -
|  | Use distance variable to get distance | X
| Speed variable path |  | TS:SR_ExternalVehicleController_RoadSpeed
| Distance variable path |  | TS:SR_ExternalVehicleController_Distance
| Simulate speed feedback |  | -
| Simulate pendant input |  | -
| Start pendant variable path |  | TS:CARBCell1.VETS_CANopen IO Subsystem_WAGO 337 4DI 4DO 16AI 4AO 2UC_DI01
| Run pendant variable path |  | TS:CARBCell1.VETS_CANopen IO Subsystem_WAGO 337 4DI 4DO 16AI 4AO 2UC_DI02
| Stop pendant variable path |  | TS:CARBCell1.VETS_CANopen IO Subsystem_WAGO 337 4DI 4DO 16AI 4AO 2UC_DI03
| Hold pendant variable path |  | TS:CARBCell1.VETS_CANopen IO Subsystem_WAGO 337 4DI 4DO 16AI 4AO 2UC_DI02
| Start pendant to detect rising edge |  | X
| Run pendant to detect rising edge |  | X
| Stop pendant to detect rising edge |  | X
| Hold pendant to detect rising edge |  | X
Connection | Interface Address
| Command Port |  | 32001
| Data Port |  | 32000
Use dedicated RTX NIC for UDP data
I/O | Channel #01 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI01
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI01
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #02 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI02
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI02
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #03 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI03
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI03
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #04 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI04
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI04
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #05 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI05
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI05
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #06 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI06
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI06
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #07 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI07
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI07
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #08 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI08
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI08
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #09 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI09
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI09
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #10 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI10
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI10
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #11 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI11
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI11
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #12 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI12
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI12
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #13 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI13
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI13
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #14 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI14
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI14
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #15 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI15
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI15
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #16 | Type | Analog Input (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AI16
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AI16
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #17 | Type | Analog Output (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AO01
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AO01
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #18 | Type | Analog Output (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AO02
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AO02
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #19 | Type | Analog Output (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AO03
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AO03
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #20 | Type | Analog Output (-10 … +10 VDC; 0 … 10 VDC)
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC AO04
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC AO04
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #21 (used by STARS VETS) | Type | Up Counter/Enable Input
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC UC01
|  | Name | Dyno Pulse Counter
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #22 (used by STARS VETS) | Type | Frequency Counter 0.1 Hz - 100 kH
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC UC02
|  | Name | Dyno Pulse Frequency
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #23 (used by STARS VETS) | Type | Digital Input Module DC 24 V 2- to 3-conductor connection; high-side switching
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC DI01
|  | Name | Start Pendant
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #24 (used by STARS VETS) | Type | Digital Input Module DC 24 V 2- to 3-conductor connection; high-side switching
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC DI02
|  | Name | Run/Hold Pendant
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #25 (used by STARS VETS) | Type | Digital Input Module DC 24 V 2- to 3-conductor connection; high-side switching
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC DI03
|  | Name | Stop Pendant
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #26 | Type | Digital Input Module DC 24 V 2- to 3-conductor connection; high-side switching
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC DI04
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC DI04
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #27 | Type | Relay Output Module 125 V AC, 30 V DC
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC DO01
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC DO01
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #28 | Type | Relay Output Module 125 V AC, 30 V DC
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC DO02
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC DO02
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #29 | Type | Relay Output Module 125 V AC, 30 V DC
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC DO03
|  | Name | WAGO 337 4DI 4DO 16AI 4AO 2UC DO03
Quantity
Unit
Scale
Offset
Log
Logging Interval
| Channel #30 (used by STARS VETS) | Type | Relay Output Module 125 V AC, 30 V DC
|  | ID | WAGO 337 4DI 4DO 16AI 4AO 2UC DO04
|  | Name | FanControl
Quantity
Unit
Scale
Offset
Log
Logging Interval

## CVS
CVS-7200 | Flow Rates | Main Venturis | Number of flow rates | 1 |  | -
|  |  |  | 2 |  | - | Has clean and dirty line
|  |  |  | 3 |  | X | dirty for phase 1
|  |  |  | ... |  | -
|  |  |  | n |  | -
|  |  | List Flow rate (m3/min) | 1 |  | 3
|  |  |  | 2 |  | 6
|  |  |  | 3 |  | 9
|  | Bag fill | Fill bag using mass flow controller |  |  | X
|  |  | Maximum sample flow rate (l/min) |  |  | ?
|  |  | Minimum sample flow rate (l/min) |  |  | ?
|  |  | Do not set bag sample flow to CVS |  |  | -
|  | Sample Venturis | Number of flow rates | 1 |  | -
|  |  |  | 2 |  | -
|  |  |  | 3 |  | -
|  |  |  | ... |  | -
|  |  |  | n |  | -
|  |  | List | Number |  | -
|  |  |  | Flow rate (l/min) |  | -
|  | Controller Type | 7x00 Series MCU |  |  | -
|  |  | ONE Series DMC |  |  | X
|  | Stop to switch flow |  |  |  | -
|  | Reference temperature (degC) |  |  |  | 20
|  | Line leak check |  |  |  | ?
| Bag Configuration | Bag Sets | Number of bag sets | 1 |  | X
|  |  |  | 2 |  | -
|  |  |  | 3 |  | -
|  |  |  | 4 |  | -
|  |  |  | 5 |  | -
|  |  | List | Number |  | 1
|  |  |  | Name | STD | -
[5 empty] | LEV | X
[5 empty] | DIESEL | -
|  |  |  | Line name | Selected
Name
|  | Max bag pairs | 1 |  |  | -
|  |  | 2 |  |  | -
|  |  | 3 |  |  | -
|  |  | 4 |  |  | -
|  |  | 5 |  |  | X
|  |  | 6 |  |  | -
|  | Bag volume (l) |  |  |  | 90
|  | Max bag fill (%) |  |  |  | 90
|  | Individual bag leak checks |  |  |  | ?
|  | Individual bag dump |  |  |  | ?
| SAO Offsets | Dilution air SAO |  |  |  | -
|  | Exhaust shut-off valve(s) |  |  |  | -
|  | SAO offset test type | Standalone |  |  | -
|  |  | Pre-test |  |  | -
|  |  | None |  |  | -
| Connection | IP address
TCP port
UDP port
CVS ONE | Flow Rates | Main Venturis | Number of flow rates | 1 |  | -
|  |  |  | 2 |  | -
|  |  |  | 3 |  | X
|  |  |  | ... |  | -
|  |  |  | n |  | -
|  |  | List Flow rate (m3/min) | 1 |  | 3
|  |  |  | 2 |  | 6
|  |  |  | 3 |  | 9
|  | Bag fill | Fill bag using mass flow controller |  |  | -
|  |  | Maximum sample flow rate (l/min) |  |  | -
|  |  | Minimum sample flow rate (l/min) |  |  | -
|  |  | Do not set bag sample flow to CVS |  |  | -
|  | Sample Venturis | Number of flow rates | 1 |  | ?
|  |  |  | 2 |  | ?
|  |  |  | 3 |  | ?
|  |  |  | ... |  | ?
|  |  |  | n |  | ?
|  |  | List | Number |  | ?
|  |  |  | Flow rate (l/min) |  | ?
|  | Controller Type | 7x00 Series MCU
|  |  | ONE Series DMC |  |  | X
|  | Stop to switch flow |  |  |  | -
|  | Reference temperature (degC) |  |  |  | 20
|  | Line leak check |  |  |  | ?
| Bag Configuration | Bag Sets | Number of bag sets | 1 |  | X
|  |  |  | 2 |  | -
|  |  |  | 3 |  | -
|  |  |  | 4 |  | -
|  |  |  | 5 |  | -
|  |  | List | Number |  | 1
|  |  |  | Name | STD | X
[5 empty] | LEV | -
[5 empty] | DIESEL | -
|  |  |  | Line name | Selected | ?
[5 empty] | Name | ?
|  | Max bag pairs | 1 |  |  | -
|  |  | 2 |  |  | -
|  |  | 3 |  |  | -
|  |  | 4 |  |  | X
|  |  | 5 |  |  | -
|  |  | 6 |  |  | -
|  | Bag volume (l) |  |  |  | 90
|  | Max bag fill (%) |  |  |  | 90
|  | Individual bag leak checks |  |  |  | ?
|  | Individual bag dump |  |  |  | ?
| SAO Offsets | Dilution air SAO |  |  |  | -
|  | Exhaust shut-off valve(s) |  |  |  | -
|  | SAO offset test type | Standalone |  |  | -
|  |  | Pre-test |  |  | -
|  |  | None |  |  | -
| Connection | IP address
TCP port
UDP port

## Flow Streams
Name |  | Standard
Valid bag sets | STD | X
| LEV | -
| DIESEL | -
Type | Tunnel | X
| RMT | -
| RawOnly | -
Dilution air SAO
Valid sample lines
Name |  | SULEV
Valid bag sets | STD | -
| LEV | X
| DIESEL | -
Type | Tunnel | X
| RMT | -
| RawOnly | -
Dilution air SAO
Valid sample lines

## Weather Station
Humidity measurement | RelativeHumidity | X
| DewPoint | -
| WetDry | -
Default cell temp (degC) |  | 20
Default barometer (kPa) |  | 101.325
Default relative humidity (%) |  | 40
Default dew point temp (degC) |  | 10
Default wet bulb temp (degC) |  | 10
Dilution air humidity measurement | RelativeHumidity | -
| DewPoint | -
| WetDry | -
| FromCellAir | X
| Fixed | -
Fixed dilution air humidity (g/kg) |  | 5.8

## Configurations Overview
| MAHA 4WD Dyno | MEXA-ONE-C1 (Bag or Contin) | CVS-ONE (STANDARD) | MEXA-ONE-C1-SL (Bag) | CVS-7200 (SULEV) | MEXA-1100QL-N2O (Bag or Contin) | MEXA-1400QL-NX NO, NO2, N2O, NH3 (Direct) | GC-Sampler | Impinger | Quad-PM | MEXA-2000SPCS Particle Count | TSI UPC | TSI EEPS | AVL 483 Soot | CO(L) AIA-11 | CO(H)/CO2 AIA-33 | NOx CLA-01 | CH4 GFA-01 | THC FIA-01 | H.THC FIA-01-O | 1100QL-N2O | CO(L) AIA-11 | CO2 (L) AIA-32 | NOx CLA-01SL | CH4 GFA-01SL | THC FIA-01SL | O2 MPA-01 | Ambient Bag 1 A1 | Ambient Bag 2 A2 | Ambient Bag 3 A3 | Ambient Bag 4 A4 | Sample Bag 1 S1 | Sample Bag 2 S2 | Sample Bag 3 S3 | Sample Bag 4 S4 | Ambient Bag 1 A1 | Ambient Bag 2 A2 | Ambient Bag 3 A3 | Ambient Bag 4 A4 | Ambient Bag 5 A5 | Sample Bag 10 S11 | Sample Bag 6 S13 | Sample Bag 7 S14 | Sample Bag 8 S14 | Sample Bag 9 S14
Gasoline Bag (Standard) | X | X | X | - | - | O | O | O | - | O | O | O | O | O | Bag | Bag | Bag | Bag | - | Bag | Bag | - | - | - | - | - | - | Ph1/CS | X | X | X | Ph1/CS | X | X | X | - | - | - | n/a | - | - | - | - | n/a | -
Ethanol Bag (Standard) | X | X | X | - | - | O | O | O | X | O | O | O | O | O | Bag | Bag | Bag | Bag | - | Bag | Bag | - | - | - | - | - | - | Ph1/CS | X | X | X | Ph1/CS | X | X | X | - | - | - | n/a | - | - | - | - | n/a | -
Diesel Bag (Standard) | X | X | X | - | - | O | O | O | - | X | O | O | O | O | Bag | Bag | Bag | Bag | - | Integ & amb bag | Bag | - | - | - | - | - | - | Ph1/CS | X | X | X | Ph1/CS | X | X | X | - | - | - | n/a | - | - | - | - | n/a | -
Gasoline Bag (SULEV) | X | - | F | X | X | O | O | O | - | O | O | O | O | O | - | - | - | - | - | - | Bag | Bag | Bag | Bag | Bag | Bag | Bag | - | - | - | - | - | - | - | - | X | X | X | n/a | Ph1/CS | X | X | X | n/a | Ph1/CS
Gasoline Bag (Standard) + Continuous Dilute | X | X | X | - | - | O | O | O | - | O | O | O | O | O | Dil/Bag | Dil/Bag | Dil/Bag | Bag | - | Dil/Bag | Bag & Dil | - | - | - | - | - | - | Ph1/CS | X | X | X | Ph1/CS | X | X | X | - | - | - | n/a | - | - | - | - | n/a | -
Diesel Bag (Standard) + Continuous Dilute | X | X | X | - | - | O | O | O | - | X | O | O | O | O | Dil/Bag | Dil/Bag | Dil/Bag | Bag | - | Integ , Dil & amb bag | Bag & Dil | - | - | - | - | - | - | Ph1/CS | X | X | X | Ph1/CS | X | X | X | - | - | - | n/a | - | - | - | - | n/a | -
Gasoline Bag (SULEV) + Continuous Dilute | X | X | F | X | X | O | O | O | - | O | O | O | O | O | Dilute | Dilute | Dilute | Bag | - | Dilute | Bag & Dil | Bag | Bag | Bag | Bag | Bag | Bag | - | - | - | - | - | - | - | - | X | X | X | n/a | Ph1/CS | X | X | X | X | Ph1/CS
[15 empty] | MEXA-ONE-C1-OV Standard/Conti Bench | [6 empty] | MEXA-ONE-C1-SL SULEV Bag Bench | [5 empty] | CVS-ONE (STANDARD) | [7 empty] | CVS-7200 (SULEV)
X : Active
- : Not used
O : Optional
F- Flowrate Set Only
Ph1- Phase 1 of FTP
CS- Cold Start

## Filter Configurations
Particulate Sampling Configuration options |  |  |  |  | User Inputs: | [17 empty] | 1066.815 (b)(1)
Filter face velocity 100 cm/sec (for 47mm PM filter) = |  | 61 | slpm |  | Filter Configurations:
Proportional Sampling Mode |  |  |  |  | a)Primary
4 lines, 3+ 1 bypass |  |  |  |  | b)Primary + Secondary
Number of Streams: 1-4
Stream A | FTP | [11 empty] | HWFET |  |  | US06
| Phase1 |  |  | Phase 2 |  |  | Phase 3 |  |  | Phase 4 (opt, requires 4 filters) |  |  | HWFET |  |  | City |  |  | Highway
| Filter Number | Filter Flowrate (slpm) | Sec. Dilution Air Flowrate (slpm) | Filter Number | Filter Flowrate (slpm) | Sec. Dilution Air Flowrate (slpm) | Filter Number | Filter Flowrate (slpm) | Sec. Dilution Air Flowrate (slpm) | Filter Number | Filter Flowrate (slpm) | Sec. Dilution Air Flowrate (slpm) | Filter Number | Filter Flowrate (slpm) | Sec. Dilution Air Flowrate (slpm) | Filter number | Filter Flowrate (slpm) | Sec. Dilution Air Flowrate (slpm) | Filter number | Filter Flowrate (slpm) | Sec. Dilution Air Flowrate (slpm)
PM Research | FN | FR | DA | FN | FR | DA | FN | FR | DA | n/a | n/a | n/a | - | - | - | - | - | - | - | - | -
PM -Part 86 sampling | 1A | 75 | 0 | 2A | 75 | 0 | 3A | 75 | 0 | n/a | n/a | n/a | - | - | - | - | - | - | - | - | -
PM- Part 1066.815 option 1 | 1A | 61 | 0 | 2A | 61 | 0 | 3A | 61 | 0 | n/a | n/a | n/a | - | - | - | - | - | - | - | - | -
PM- Part 1066.815 option 2 | 1A | 61 | 0 | 1A | 61 | 0 | 2A | 61 | 0 | n/a | n/a | n/a | - | - | - | - | - | - | - | - | -
PM- Part 1066.815 option 3* | 1A | 61 | 0 | 1A | 61 | 0 | - | - | - | n/a | n/a | n/a | - | - | - | - | - | - | - | - | -
PM- Part 1066.815 option 4a | 1A | 26.23 | 0 | 1A | 61 | 0 | 1A | 34.77 | 0 | n/a | n/a | n/a | - | - | - | - | - | - | - | - | -
PM- Part 1066.815 option 4b | 1A | 61 | 34.77 | 1A | 61 | 0 | 1A | 61 | 26.23 | n/a | n/a | n/a | - | - | - | - | - | - | - | - | -
PM- Part 1066.815 option 5a* | 1A | 45.75 | 0 | 1A | 45.75 | 0 | 1A | 61 | 0 | n/a | n/a | n/a | - | - | - | - | - | - | - | - | -
PM- Part 1066.815 option 5b* | 1A | 61 | 15.25 | 1A | 61 | 15.25 | 1A | 61 | 0 | n/a | n/a | n/a | - | - | - | - | - | - | - | - | -
PM- HwFET | - | - | - | - | - | - | - | - | - | n/a | n/a | n/a | 1A | 61 | 0 | - | - | - | - | - | -
PM- US06 | - | - | - | - | - | - | - | - | - | n/a | n/a | n/a | - | - | - | 1A | 61 | 0 | 1A | 61 | 0 | § 1066.815 Exhaust emission test procedures for FTP testing.
PM- US06- split bag | - | - | - | - | - | - | - | - | - | n/a | n/a | n/a | - | - | - | 1A | 61 | 0 | 2A | 61 | 0 | (a) General. The FTP exhaust emission test sequence consists of a cold-start test and a hot-start
* Requires Stream A and Stream B | FN: FN= 1-3; FR: FR = 9-75; DA: DA=0 or 5-75 | test as described in § 1066.801.
(b) PM sampling options. Collect PM using any of the procedures specified in paragraphs (b)(1)
through (5) of this section and use the corresponding equation in § 1066.820 to calculate FTP
composite emissions. Testing must meet the requirements related to filter face velocity as
described in § 1066.110(b)(2)(iii)(C), except as specified in paragraphs (b)(4) and (5) of this
section. For procedures involving flow weighting, set the filter face velocity to a weighting
target of 1.0 to meet the requirements of § 1066.110(b)(2)(iii)(C). Allow filter face velocity to
decrease as a percentage of the weighting factor if the weighting factor is less than 1.0 and do not
change the nominal CVS flowrates or secondary dilution ratios between FTP or UDDS test
intervals. Use the appropriate equations in § 1066.610 to show that you meet the dilution factor
requirements of § 1066.110(b)(2)(iii)(B). If you collect PM using the procedures specified in
paragraphs (b)(4) or (5) of this section, the residence time requirements in 40 CFR
1065.140(e)(3) apply, except that you may exceed an overall residence time of 5.5 s for sample
flow rates below the highest expected flow rate.
(1) You may collect a separate PM sample for transient and stabilized portions of the cold-start
UDDS and the hot-start UDDS. This may either be done by sampling with three bags or four
bags. You may omit the stabilized portion of the hot-start test (bag 4) and use the stabilized
portion of the cold-start test (bag 2) in its place.
(2) You may collect PM on one filter over the cold-start UDDS and on a separate filter over the
(2) You may collect PM on one filter over the cold-start UDDS and on a separate filter over the
hot-start UDDS.
(3) You may collect PM on one filter over the cold-start UDDS (bag 1 and bag 2) and on a
separate filter over the 867 seconds of the stabilized portion of the cold-start UDDS and the first
505 seconds of the hot-start UDDS (bag 2 and bag 3). Note that this option involves duplicate
measurements during the stabilized portion of the cold-start UDDS.
(4) You may collect PM on a single filter over the cold-start UDDS and the first 505 seconds of
the hot-start UDDS using one of the following methods:
(i) Adjust your sampling system flow rate over the filter to weight the filter face velocity over the
three intervals of the FTP based on weighting targets of 0.43 for bag 1, 1.0 for bag 2, and 0.57
for bag 3.
(ii) Maintain a constant sampling system flow rate over the filter for all three intervals of the FTP
by increasing overall dilution ratios for bag 1 and bag 3. To do this, reduce the sample flow rate
from the exhaust (or diluted exhaust) such that the value is reduced to 43 % and 57 %,
respectively, of the bag 2 values. For constant-volume samplers, this requires that you decrease
the dilute exhaust sampling rate from the CVS and compensate for that by increasing the amount
of secondary dilution air.
(5) You may collect PM on a single filter over the cold-start UDDS and the full hot-start UDDS
using one of the following methods:
(2) You may collect PM on one filter over the cold-start UDDS and on a separate filter over the
(i) Adjust your sampling system flow rate over the filter to weight the filter face velocity based
on weighting targets of 0.75 for the cold-start UDDS and 1.0 for the hot-start UDDS.
(ii) Maintain a constant sampling system flow rate over the filter for both the cold-start and hotstart
UDDS by increasing the overall dilution ratio for the cold-start UDDS. To do this, reduce
the sample flow rate from the exhaust (or diluted exhaust) such that the value is reduced to 75 %
of the hot-start UDDS value. For constant-volume samplers, this requires that you decrease the
dilute exhaust sampling rate from the CVS and compensate for that by increasing the amount of
secondary dilution air.
(c) Gaseous sampling options. Collect gaseous samples using any of the following procedures:
(1) You may collect a single sample for a full UDDS (cold-start or hot-start).
(2) You may sample emissions separately for transient and stabilized portions of any UDDS.
(3) You may omit the stabilized portion of the hot-start test (bag 4) and use the stabilized portion
of the cold-start test (bag 2) in its place.
(d) Test sequence. Follow the exhaust emission measurement procedures specified in §§
1066.410 through 1066.425, subject to the following exceptions and additional provisions:
(1) Take the following steps for the cold-start test:
(i) Precondition the vehicle as described in § 1066.816. Initiate the cold-start test following the
12 to 36 hour soak period.
(2) You may collect PM on one filter over the cold-start UDDS and on a separate filter over the
(ii) Start sampling and recording simultaneously with starting the vehicle. Place the vehicle in
gear 15 seconds after engine starting, which is 5 seconds before the first acceleration.
(iii) At the end of the deceleration scheduled to occur 505 seconds into the cold-start UDDS,
simultaneously switch all the sample flows from the cold-start transient interval to the stabilized
interval, stopping all cold-start transient interval sampling and recording, including background
sampling. Reset integrating devices for the stabilized interval and indicate the end of the coldstart
interval in the recorded data. Operate the vehicle over the remainder of the UDDS. Turn
the engine off 2 seconds after the end of the last deceleration in the stabilized interval (1,369
seconds after the start of the driving schedule).
(iv) Five seconds after the engine stops running, stop all stabilized interval sampling and
recording, including background sampling. Stop any integrating devices for the stabilized
interval and indicate the end of the stabilized interval in the recorded data. Note that the 5
second delay is intended to account for sampling system transport.
(2) Take the following steps for the hot-start test:
(i) Initiate the hot-start test (9 to 11) minutes after the end of the sample period for the cold-start
UDDS.
(ii) Repeat the steps in paragraph (d)(1)(ii) of this section. Operate the vehicle over the first 505
seconds of the UDDS. For tests that do not include bag 4 operation, turn off the engine and
simultaneously stop all hot-start sampling and recording, including background sampling, and
any integrating devices at the end of the deceleration scheduled to occur 505 seconds into the
hot-start UDDS.
(iii) To include bag 4 measurement, operate the vehicles over the remainder of the UDDS and
conclude the testing as described in paragraphs (d)(1)(iii) and (iv) of this section.
(3) This completes the procedure for measuring FTP exhaust emissions. See § 1066.801 and
subpart J of this part for continuing the test sequence to measure evaporative or refueling
emissions.
§ 1066.816 Vehicle

## Configurations
Gasoline Bag Only (Standard) | Settings | Sample bags |  | X
|  | Bag set | STD | X
LEV
DIESEL
Certification
|  | Bag range | Sniff | X
|  |  | Auto range | -
|  |  | Fixed | -
|  | Bag analyzer calibration | None | -
|  |  | Zero/Span | -
|  |  | Zero/Span with Drift Check | X
|  | Do rezero at calibration |  | -
|  | Do rezero at drift check |  | -
|  | Flow Stream | Gasoline | X
|  |  | Diesel | -
|  | Continuous calibration | None | -
|  |  | Zero/Span | -
|  | Calibration sharing | None | -
Between bag read for adjacent phases
|  |  | Between bag reads for a test procedure | -
Bag calibration at test start
|  | Report |  | Select Report Template
|  | Bag purge pre-test |  | ?
|  | CVS line leak check pre-test |  | ?
|  | Bag dump post-test |  | ?
|  | Bag purge post-test |  | ?
|  | Purge lines pre-test |  | ?
|  | Calibrate integrated components separately during soak |  | ?
|  | Delay soak completion to allow bag analysis to complete |  | -
| Continuous Lines | Selected
Line name
|  | Component Ranges | COL
Range
COH
Range
CO2
Range
CH4
Range
| Integrated Components | Selected
Line name
Component Ranges
Note on D9: Siebler, Jorg: reads sample bag range on auto to find appropriate range and then locks that range before commencing a bag read
Note on D10: Siebler, Jorg: reads bag range on auto range
Note on D11: Siebler, Jorg: reads bag ranges as specified in Bag components and ranges fields
Note on D12: Siebler, Jorg: no calibration is performed
Note on D13: Siebler, Jorg: performs a zero/span adjustment
Note on C15: Siebler, Jorg: performs a rezero adjustment when the bag analyzer calibration is done
Note on C16: Siebler, Jorg: performs a rezero adjustment when the drift check is done
Note on D20: Siebler, Jorg: performs a device-specific Zero Span
Note on C21: Siebler, Jorg: enables the calibration test specified in the Bag analyzer calibration field to be run for either adjacent phases or for a test procedure. Options are; None, Between bag reads for adjacent phases (only enabled if Fixed is selected in the Bag Range field) or Between bag reads for a Test Procedure
Note on C24: Siebler, Jorg: If enabled and selected, bag calibration is done during the active driving part of the test if the selected analyzers are only used for bag read. It is done in the pre-test sequence if the selected analyzers are used for continuous measurement
Note on C26: Siebler, Jorg: Select if the bag sets need to be purged prior to testing
Note on C27: Siebler, Jorg: Select if the bag sets need to be purged prior to testing
Note on C28: Siebler, Jorg: Select if the bag sets need to be 'dumped' after testing
Note on C29: Siebler, Jorg: Select if the bag sets need to be purged after testing
Note on C30: Siebler, Jorg: Select to confirm that the Continuous Lines are purged before calibration for a duration specified in the Global Settings
Note on C31: Siebler, Jorg: calibration of integrated components is executed separately, after the bag analysis for normal components during soak, if the remaining soak time is enough for the calibration
Note on C32: Siebler, Jorg: completion of the soak state is blocked until the bag analysis is completed
Note on B35: Siebler, Jorg: Needs to be adjusted

## Optional Device Config.
SPCS | Sample SPCS |  | STARS VETS Variable
PND1 to set
PND2 to set
CPC/CFO Flow Cal. and Diff Pressure Cal.
CPC status check
System status check
NMOG Sampler | Use NMOG sampler |  | STARS VETS Variable
Leak check
| Flow rate mode | Fixed flow rate (l/min)
Proportional flow (l/min)
Sample alcohol
Sample aldehyde
MEXA 1400 | Use Device |  | STARS VETS Variable
Pre Test Calibrate
Post Test Purge
Purge Time
Purge A
Purge B
Note on B6: Siebler, Jorg: defines the primary dilution factor to be used and can not be blank if Sample SPCS is selected
Note on B7: Siebler, Jorg: defines the secondary dilution factor to be used and can not be blank if Sample SPCS is selected
Note on B8: Siebler, Jorg: if selected, the CPC/CFO Flow Calibration and Diff. Pressure Calibration are executed on the SPCS
Note on B9: Siebler, Jorg: if selected a status check of the counter is performed at the start of the test
Note on B10: Siebler, Jorg: if selected a status check of the counter is performed at the start of the test
Note on B12: Siebler, Jorg: Use the NMOG Sampler for the specified test configuration
Note on B13: Siebler, Jorg: Check for leaks during the pre-test procedure. If a leak is detected a dialog will prompt on how you can proceed
Note on B16: Siebler, Jorg: To detect alcohol compounds
Note on B17: Siebler, Jorg: To detect aldehyde compounds

## Settings
Results Processing | Report name
Exporter
| Name format | Test name
Fuel name
Daily Test No.
Time
Test procedure name
Vehicle name
Date
Custom Field
Example
Bag Fill Settings | When calculating bag sample rate use | Calculated bag fill with margin |  | -
|  | Bag fill fraction |  | X
| Volume margin above that required for bag read (l) |  |  | -
| Target bag fill (%) |  |  | 90
Bag Read Times | Measurement Type | Default | Sniff read delay (s) | 5
|  |  | Sample bag read delay (s) | 10
|  |  | Ambient bag read delay (s) | 10
|  |  | Auto range time (s) | 5
|  |  | Bag read time window (s) | 5
|  |  | Maximum bag read time (s) | 30
|  |  | Target stability ratio | 0.03
|  |  | Zero time (s) | 10
|  |  | Span time (s) | 10
|  |  | Second zero time (s) | 10
|  |  | Drift check tolerance | 0.02
|  | LE | Sniff read delay (s) | 30
|  |  | Sample bag read delay (s) | 40
|  |  | Ambient bag read delay (s) | 40
|  |  | Auto range time (s) | 10
|  |  | Bag read time window (s) | 10
|  |  | Maximum bag read time (s) | 30
|  |  | Target stability ratio | 0.03
|  |  | Zero time (s) | 40
|  |  | Span time (s) | 40
|  |  | Second zero time (s) | 40
|  |  | Drift check tolerance | 0.02
|  | GC | Sniff read delay (s) | 5
|  |  | Sample bag read delay (s) | 10
|  |  | Ambient bag read delay (s) | 10
|  |  | Auto range time (s) | 5
|  |  | Bag read time window (s) | 5
|  |  | Maximum bag read time (s) | 30
|  |  | Target stability ratio | 0.03
|  |  | Zero time (s) | 10
|  |  | Span time (s) | 10
|  |  | Second zero time (s) | 10
|  |  | Drift check tolerance | 0.02
|  | GC-LE | Sniff read delay (s) | 30
|  |  | Sample bag read delay (s) | 50
|  |  | Ambient bag read delay (s) | 50
|  |  | Auto range time (s) | 20
|  |  | Bag read time window (s) | 10
|  |  | Maximum bag read time (s) | 30
|  |  | Target stability ratio | 0.03
|  |  | Zero time (s) | 50
|  |  | Span time (s) | 50
|  |  | Second zero time (s) | 50
|  |  | Drift check tolerance | 0.02
|  | QCL | Sniff read delay (s) | 5
|  |  | Sample bag read delay (s) | 10
|  |  | Ambient bag read delay (s) | 10
|  |  | Auto range time (s) | 5
|  |  | Bag read time window (s) | 5
|  |  | Maximum bag read time (s) | 30
|  |  | Target stability ratio | 0.03
|  |  | Zero time (s) | 10
|  |  | Span time (s) | 10
|  |  | Second zero time (s) | 10
|  |  | Drift check tolerance | 0.02
Exhaust flow method | Diesel | None |  | -
|  | Dilution Air Flow |  | X
|  | Tracer |  | -
|  | Measured |  | -
|  | Dilution Air Flow Tracer Hybrid |  | -
| Gasoline | None |  | -
|  | Dilution Air Flow |  | X
|  | Tracer |  | -
|  | Measured |  | -
|  | Dilution Air Flow Tracer Hybrid |  | -
NMOG Settings | Type of data provided by the chemical laboratory | Mass in mg |  | X
|  | Concentration in ppm |  | -
|  | None |  | -
| Manual entry volume reference temperature (degC) |  |  | 20
Particulate Settings | Send flow settings to particulate sampler |  |  | X
| Default background concentration (mg/m3) |  |  | 0
| Manage reference filters |  |  | -
| Apply buoyancy correction analysis |  |  | X
| Calibration weight density (kg/m3) |  |  | 0
| Particulate filter density (kg/m3) |  |  | 0
| Default absolute pressure during weighing (kPa) |  |  | 101.325
| Default temperature during weighing (degC) |  |  | 22
Global Settings | Clear messages on test start
| Stop CVS blower in soak phase |  |  | X
Prompt for test start
| Continuous line purge time (s) |  |  | 30
Run analysis on test abort
| Continuous data logging | MeasurementPhasesOnly |  | X
StartAndDrivePeriods
StartDriveAndSoakPeriods
EntireTest
| Prompt in case of calibration failure |  |  | X
| Calibration failure dialog timeout (s) |  |  | 60
Allow test with no CVS blower running
| Wait for run timeout (s) |  |  | 30
| Wait for stop timeout (s) |  |  | 30
| Background validation interval (min) |  |  | 0
Check road load model match before test
Proportional PM flow sampling
| Calibrate SAO pressure before test |  |  | X
| Stop CVS blower before SAO pressure calibration |  |  | X
Send SAO offset to device
Proportional sample bag sampling
Proportional ambient bag sampling
Tailpipe pressure control
| Tailpipe pressure (Pa) |  |  | 0
| CVS Blower Stop Dialog Timeout (s) |  |  | 30
| QCL zero/span mode | Bag |  | X
Continuous
BagAndContinuous
Read sample bag before ambient bag
Note on B108: Siebler, Jorg: all previous error messages, logs, etc. will be cleared whenever a test starts
Note on B109: Siebler, Jorg: the CVS Blower will be stopped when any test cycle enters a Soak Phase
Note on B110: Siebler, Jorg: a prompt displays when a resource is initialized
Note on B111: Siebler, Jorg: specifies the purge duration for Continuous Lines prior to calibration, whenever a MEXA device is initialized
Note on B112: Siebler, Jorg: an analysis is still performed using the data that was logged prior to the test being aborted
Note on B117: Siebler, Jorg: causes an action prompt asking the user to ignore or retry the calibration is displayed in the event of a MEXA calibration failure
Note on B118: Siebler, Jorg: amend the timeout period for the MEXA calibration failure action prompt described above if required. If this dialog times out, then a notification dialog is displayed, which will remain displayed until the next test. At the end of the test, a dialog is displayed prompting the user to dump/purge the bags if there were calibration failures for the bag read
Note on B119: Siebler, Jorg: must be ticked before tests can be run with no CVS Blower in operation (must be implicitly set as there could be issues with exhaust ventilation)
Note on B120: Siebler, Jorg: amend the timeout period if required. If the timeout is exceeded after ‘start’ then the test goes to the wait start state and any bag fill is stopped. 3 tries are made after which the test will abort
Note on B121: Siebler, Jorg: amend the timeout period if required. The Wait for stop timeout(s) means that if the stop button is not pressed within the configured timeout from the end of the trace then the bag fill is stopped (thus simulating the Stop button press)
Note on B123: Siebler, Jorg: the road load model selected on the dyno is cross checked at the pre-test stage with a road load model identifier entered into the VETS system
Note on B125: Siebler, Jorg: the SAO pressure transducer zero operation is run during pre-test. This setting controls the pressure transducer zero operation for the CVS or DTSP device that is managing SAO
Note on B126: Siebler, Jorg: the CVS blower is stopped before the SAO pressure transducer is zero'd. If unticked, the SAO pressure transducer can be zero'd when the blower is running
Note on B127: Siebler, Jorg: the selected offset tables are sent to the DTSP so the flow can be corrected at the DTSP and not in the analysis
Note on B128: Siebler, Jorg: the DTSP takes a sample proportional to the CVS flow rate
Note on B129: Siebler, Jorg: the DTSP takes an ambient dilution air sample at a rate proportional to the CVS
Note on B130: Siebler, Jorg: the DTSP device is used to control tail pipe pressure and the Tail pipe pressure (Pa) field is enabled
Note on B136: Siebler, Jorg: determines the bag read order. This defaults to false. The variable is set on the MEXA driver in MEXADevice

## Sample Line Configurations
Name | Line name
Sample Point
Delay (s)